## Explanation
Table 1 | Column | Description
| J | Enter model and serial number of your instrument where indicated above data table.
| A | Enter the actual value of your specific calibration fluid. This value can be found on the bottle label or on the accompanying calibration certificate.
| B | No entry required. Calculated 1% of fluid value in column A
| C | No entry required. Shows instrument model.
| D | Enter spindle code for spindle used in calibration check
| E | Enter RPM used for calibration check
| F | No entry required, FSR is calculated based upon spring, spindle and speed
| G | No entry required. Calculated 1% of FSR value in column F
| H | No entry required. % torque value calculated from viscosity in column I
| I | Enter viscosity reading from instrument display
Table 2 | Column | Description
| Low Limit | No entry required. Calculated value for low limit of acceptable range
| Reading | No entry required. Value copied from column I in table 1
| High Limit | No entry required. Calculated value for high limit of acceptable range
| Passed? | TRUE = Passed. FALSE = Failed

## Reference Only
Sp Code | SMC |  |  |  | Torque | TK
0 | 0.64 | ULA
1 | 1 |  | Spindle selected | SMC | 0.25RV | 0.25
2 | 4 |  | 27 | 25 | 0.5RV | 0.5
3 | 10 | RV/HA/HB |  |  | 2.5HA | 5
4 | 20 | spindles | Spring Selected | TK | 2.5HB | 20
5 | 40 |  | LV | 0.09373 | 2.5LV | 0.23432
6 | 100 |  |  |  | 2HA | 4
7 | 400 |  |  |  | 5HB | 40
14 | 125 |  |  |  | 5LV | 0.46865
15 | 50 |  |  |  | HA | 2
16 | 128 |  |  |  | HB | 8
18 | 3.2 | Small |  |  | LV | 0.09373
21 | 5 | Sample |  |  | RV | 1
25 | 512 | Adapter
27 | 25
28 | 50
29 | 100
31 | 32
34 | 64
37 | 25
40 | 0.327
41 | 1.228 | W-B
42 | 0.64 | Cone
51 | 5.178 | spindles
52 | 9.922
61 | 6.4
62 | 32
63 | 128 | LV
64 | 640 | spindles
65 | 1280
66 | 32
67 | 128
70 | 105
71 | 2.62
72 | 11.1 | Vane
73 | 53.5 | spindles
74 | 535
81 | 3.7
82 | 3.75
83 | 12.09 | DIN
85 | 1.22 | spindles
86 | 3.65
87 | 12.13
91 | 20
92 | 40
93 | 100 | Helipath
94 | 200 | spindles
95 | 500
96 | 1000

## Calibration Data
CALIBRATION TEMPLATE FOR ANALYZING COAXIAL CYLINDER CALIBRATION RESULTS
THIS CALIBRATION TEMPLATE TAKES INTO ACCOUNT THE 2% ACCURACY OF COAXIAL CYLINDER GEOMETRIES
| INSTRUCTIONS : | ENTER VALUES IN |  | BOLD | FIELDS ONLY
Enter value of calibration fluid in column A | [6 empty] | Enter instrument model: | LV
^ | [6 empty] | serial number:
¦ |  |  | Enter spindle code in column D
¦ |  |  | ^ | Enter in numerical order the RPMs used In column E
¦ |  |  | ¦ | ^
¦ |  |  | ¦ | ¦ |  |  |  |  | Enter viscosity reading in column I
A | B | C | D | E | F | G | H | I
Actual Fluid Value (cP) from Label or Certificate | 1% of Fluid Value (cP) | Instrument Model | Spindle Code | RPM | FSR cP | 2% of FSR cP | % Torque (must be >10%) | Viscosity Reading in cP
498 | 4.98 | LV | 27 | 12 | 1952.7083 | 39.0542 | 25.5% | 498
498 | 4.98 | LV | 27 | 30 | 781.0833 | 15.6217 | 63.76% | 498
498 | 4.98 | LV | 27 | 40 | 585.8125 | 11.7163 | 85.01% | 498
498 | 4.98 | LV | 27 | 30 | 781.0833 | 15.6217 | 63.76% | 498
498 | 4.98 | LV | 27 | 12 | 1952.7083 | 39.0542 | 25.5% | 498
|  |  |  | RPM | Low Limit | Reading | High Limit | Calibration
[5 empty] | = A – (B + G) | I | = A + (B + G) | Passed?
|  |  |  | 12 | 453.9658 | 498 | 542.0342 | TRUE
|  |  |  | 30 | 477.3983 | 498 | 518.6017 | TRUE
|  |  |  | 40 | 481.3037 | 498 | 514.6962 | TRUE
NOTE: This calibration template is a method of checking calibration only. Regardless of the results obtained from this template, we recommend that all Laboratory Viscometers/Rheometers/Texture Analyzers be returned to Brookfield or an authorized dealer on an annual basis for our Calibration & Certification Service. Having this annual maintenance done will ensure your instrument’s long life and optimal performance.

## Troubleshooting
If calibration check fails, what should you check?
1) | Is temperature of fluid equilibrated at 25C? | Calibration fluid must be at specified temperature +/- 0.2C
2) | Is fluid being measured in a 600ml beaker? | Smaller containers will increase apparent viscosity
3) | Is instrument guard leg being used? (applies only to standard spindles with LV and RV models) | Without guardleg in place apparent viscosity will be lower than expected. Effect is greatest when using larger spindles.
4) | Is spindle immersion mark being observed? | Surface of fluid should be in center of immersion groove. On spindles with wide groove the immersion mark is raised ring at center of groove.
5) | Did spindle rotate long enough? (minimum of 5 revolutions) | LV model instruments at high rpm may require as long as 45 seconds to stabilize.
6) | Is the viscometer level? | Bubble level must be centered.
7) | With no spindle attached and the motor off, does the % display read 0.0? (-0.1 to +0.1 is OK) Does not apply to DV-E. | Verifys that good zero torque is established and bearing mechanism is in good condition.
8) | With no spindle attached and the motor running at 10rpm or 12rpm does the % display still show 0.0? | Further verification that good zero torque is established and bearing mechanism is in good condition.
If calibration check still fails you may have to send your instrument to Brookfield for repair.
AMETEK Brookfield
11 Commerce Blvd
Middleboro, MA 02346
Tel: 508 946-6200
Fax: 508 946-6262