## CT3 CS
CT3 TEXTURE ANALYZER
| Calibration Check Using Bloom Test Strip |  |  |  |  | serial # |  |  |  | Load Cell Accuracy
| Position | 1 | 2 | 3 | 4 | BS219 |  |  |  | 5 | g
| Cert value | 80.87 | 239.5 | 392.2 | 534.2 |  |  | all values are in grams
| Std Dev | 0.12 | 0.22 | 0.27 | 0.21
[7 empty] | Low |  | High
[7 empty] | Limit | Reading | Limit
[6 empty] | 0.4493% | 80.5067 | 80.87 | 81.2333 |  | 0.4493%
|  |  | Load Cell Capacity | 1000 | G | 0.2744% | 238.8427 | 239.5 | 240.1573 |  | 0.2744%
| Position | 1 | 2 | 3 | 4 | 0.2052% | 391.395 | 392.2 | 393.005 |  | 0.2052%
| Reading 1 | 81 | 239.4 | 392.4 | 534.3 | 0.1178% | 533.5707 | 534.2 | 534.8293 |  | 0.1178%
| Reading 2 | 80.9 | 239.6 | 392.4 | 534.4
| Reading 3 | 81 | 239.3 | 392.1 | 534.2
| Reading 4 | 80.7 | 239.4 | 392.5 | 534.3 | 1 | The mean of six readings at each support position
| Reading 5 | 80.8 | 239.4 | 391.9 | 533.8 |  | should fall between the low and high limits shown above.
| Reading 6 | 80.8 | 239.9 | 391.9 | 534.2
| MIN | 80.7 | 239.3 | 391.9 | 533.8 | 2 | This applies to all bloom test strips that have been
| MAX | 81 | 239.9 | 392.5 | 534.4 |  | certified on the CT3 with which they are being used.
| MEAN | 80.8667 | 239.5 | 392.2 | 534.2
| STD DEV | 0.12111 | 0.21909 | 0.26833 | 0.20976 | 3 | If the bloom test strip you are using has not been
| RATIO | 0.1498% | 0.09148% | 0.06842% | 0.03927% |  | certified on the CT3 with which it is being used, 20% of
the load cell accuracy must be subtracted from the low limit
and added to the high limit in the table above in order
to determine acceptable pass/fail limits.
INSTRUCTIONS FOR USE
MAKE ENTRIES IN ALL CELLS WITH GREY BACKGROUND
ALL OTHER CELLS ARE CALCULATED
Note on E10: : ENTER LOAD CELL CAPACITY IN GRAMS
Note on C5: : ENTER LOAD VALUE AS STATED ON CERTIFICATE FOR POSITION 1
Note on D5: : ENTER LOAD VALUE AS STATED ON CERTIFICATE FOR POSITION 2
Note on E5: : ENTER LOAD VALUE AS STATED ON CERTIFICATE FOR POSITION 3
Note on F5: : ENTER LOAD VALUE AS STATED ON CERTIFICATE FOR POSITION 4
Note on F12: Take six readings at position 4 on the calibration strip. Record all six results in this column
Note on E12: Take six readings at position 3 on the calibration strip. Record all six results in this column
Note on D12: Take six readings at position 2 on the calibration strip. Record all six results in this column
Note on C12: Take six readings at position 1 on the calibration strip. Record all six results in this column
Note on B22: LenT: All ratio calculations less than 1% indicate acceptable repeatability
Note on C6: LenT: ENTER STD DEV VALUE AS STATED ON CERTIFICATE FOR POSITION 1
Note on D6: LenT: ENTER STD DEV VALUE AS STATED ON CERTIFICATE FOR POSITION 2
Note on E6: LenT: ENTER STD DEV VALUE AS STATED ON CERTIFICATE FOR POSITION 3
Note on F6: LenT: ENTER STD DEV VALUE AS STATED ON CERTIFICATE FOR POSITION 4
Note on G4: LenT: ENTER SERIAL NUMBER OF BLOOM TEST STRIP

## CT3 CW
Calibration Spreadsheet
CT3 TEXTURE ANALYZER |  |  | serial # | 8484734
|  |  | Load Cell Capacity | 1000 | [5 empty] | Load Cell Accuracy
| # | 1 | 2 | 3 | 4 | 5 |  |  |  | 5 | g
| Mass | 99.79 | 149.2 | 249.01 | 498.7 |  |  | all weights are in grams
| Serial # | 421 | 388 | 258 | 369
[7 empty] | Low Limit |  | High Limit
|  |  | Weight |  |  |  |  | Reading |  | Difference |  | 0
|  | 1 | 99.79 |  |  |  | 94.79 | 99.8 | 104.79 | 0.01 |  | 0
|  | 1 & 2 | 248.99 |  |  |  | 243.99 | 248.7 | 253.99 | -0.29 |  | 0
|  | 1,2 & 3 | 498 |  |  |  | 493 | 498.1 | 503 | 0.1 |  | 0
|  | 1, 2, 3 & 4 | 996.7 |  |  |  | 991.7 | 995.5 | 1001.7 | -1.2 |  | 0
|  | 1 - 5 | 996.7 |  |  |  | 991.7 | 995.8 | 1001.7 | -0.9
INSTRUCTIONS FOR USE
MAKE ENTRIES IN ALL CELLS CONTAINING COMMENTS
FIRST DECIDE UPON THE ORDER IN WHICH YOU WANT TO ADD THE WEIGHTS, THEN ENTER THEIR MASS AND SERIAL NUMBER IN THAT ORDER
AS THE WEIGHTS ARE ADDED ENTER THE LOAD READING WHERE INDICATED
Note on E4: : ENTER LOAD CELL CAPACITY IN GRAMS
Note on C6: : ENTER MASS AS STATED ON CERTIFICATE FOR WEIGHT WITH HANGER
Note on C7: : ENTER SERIAL NUMBER OF WEIGHT WITH HANGER
Note on D6: : ENTER MASS AS STATED ON CERTIFICATE OF SECOND WEIGHT YOU WILL ADD
Note on D7: : ENTER SERIAL NUMBER OF SECOND WEIGHT
Note on E6: : ENTER MASS AS STATED ON CERTIFICATE OF THE THIRD WEIGHT YOU WILL ADD
Note on E7: : ENTER SERIAL NUMBER OF THIRD WEIGHT
Note on F6: : ENTER MASS FROM CERTIFICATE OF FOURTH WEIGHT YOU WILL ADD
Note on F7: : ENTER SERIAL NUMBER OF FOURTH WEIGHT
Note on G6: : ENTER MASS AS STATED ON CERTIFICATE OF THE FIFTH WEIGHT (IF ANY) YOU WILL ADD
Note on G7: : ENTER SERIAL NUMBER OF FIFTH WEIGHT
Note on I10: : ENTER LFRA LOAD VALUE FOR FIRST WEIGHT
Note on I11: : ENTER LFRA LOAD VALUE OF BOTH WEIGHTS
Note on I12: : ENTER LFRA LOAD VALUE OF ALL THREE WEIGHTS
Note on I13: : ENTER LFRA LOAD VALUE OF ALL FOUR WEIGHTS
Note on I14: : ENTER LFRA LOAD VALUE OF ALL FIVE WEIGHTS
Note on D16: :MAKE ENTRIES IN ALL CELLS CONTAINING COMMENTS FIRST DECIDE UPON THE ORDER IN WHICH YOU WANT TO ADD THE WEIGHTS, THEN ENTER THEIR MASS AND SERIAL NUMBER IN THAT ORDER AS THE WEIGHTS ARE ADDED ENTER THE LOAD READING WHERE INDICATED
Note on E2: : ENTER THE SERIAL NUMBER OF YOUR INSTRUMENT